Central Near-Vision (CN) Design Ortho-K Lens

R&D Test Report

Objective: Determine whether the CN-design lens delivers meaningfully better near vision than a standard Ortho-K lens, and identify the parameter sensitivities that drive the near-vision response.

1. Test Objectives

- Verify whether the CN-design lens improves near vision compared to a standard Ortho-K lens.

- Change only one parameter per iteration, and compare the four-way response: distance-with-lens, near-with-lens, distance-post-removal, near-post-removal.

2. Test Subject

| Eye | Myopia Level | Spherical Equivalent | Dominant Eye |
| --- | --- | --- | --- |
| Right (OD) | Low myopia | ≈ −2.00 D | Yes |
| Left (OS) | High myopia | ≈ −5.50 D | No |

Each eye was tuned independently, forming two parallel iteration tracks:

OD → Low-Myopia Track  ·  OS → High-Myopia Track  ·  5 rounds each.

3. Method

Baseline = prior distance-only design parameters. CN lenses were fabricated; in each round one variable was changed (C factor / ADD / Power). Four outcomes were recorded per round:

- Distance vision with lens on (on-lens distance)

- Near vision with lens on (on-lens near)

- Distance vision after removal (post-removal distance)

- Near vision after removal (post-removal near)

4. OD Iteration Track — Low Myopia (≈ −2.00 D)

| # | BC | Power | Dia. | ADD | C factor | Change | On-lens Dist. | On-lens Near | Post-removal Dist. | Post-removal Near |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| R1 | 44.25 | −1.25 | 10.6 | +1.5D | 4.0 | Baseline | Very good | Slightly better than prior MCOK | Very good | Baseline; near below expectation |
| R2 | 44.25 | −1.25 | 10.6 | +1.5D | 6.0 ↑ | C ↑ | No effect | — | Better (prior MCOK over-corrected) | C↑ no negative impact on distance; near ~same as R1 |
| R3 | 44.25 | −1.25 | 10.6 | +2.5D ↑ | 6.0 | ADD ↑ | Clear | Holding up | 0.8 (next-day wear) | ADD↑ improved near (next-day wear) |
| R4 | 44.25 | −1.25 | 10.6 | +2.5D | 4.0 ↓ | C ↓ | Clear | Similar to R3 | — | C 6.0 → 4.0: negligible difference |
| R5 | 44.25 | −1.25 | 10.6 | +2.5D | 1.0 ↓↓ | C ↓↓ | Mediocre | Worse | — | C too low; near degrades |

OD Summary

- Toggling C factor between 4.0 ↔ 6.0 produced little observable difference → low myopia appears C-factor-insensitive within this range.

- Pushing C factor to the extreme (1.0) clearly worsened near → a lower bound exists.

- Raising ADD from +1.5 D → +2.5 D improved near on next-day wear; usable at normal reading distance but not equivalent to +2.5 D reading glasses.

- Overall, near vision was modestly better than a standard Ortho-K lens, but did not reach the target outcome.

5. OS Iteration Track — High Myopia (≈ −5.50 D)

| # | BC | Power | Dia. | ADD | C factor | Change | On-lens Dist. | On-lens Near | Post-removal Dist. | Post-removal Near |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| L1 | 44.00 | −4.75 | 10.6 | +2.5D | 4.0 | Baseline | 0.6 | Comfortable | 0.6 (under-corrected) | Affected by sleep and wear duration |
| L2 | 44.00 | −6.00 ↑↑ | 10.6 | +2.5D | 4.0 | Power ↑↑ | Blurry (small-ring) | Poor | 1.2 | Multi-focal sensation at near |
| L3 | 44.00 | −5.25 ↓ | 10.6 | +2.5D | 4.0 | Power ↓ | Clear | Poor | 0.8 (Continue 2 nights wearing) | Dist. clear after down-titration; near mediocre |
| L4 | 44.00 | −5.25 | 10.6 | +2.5D | 6.0 ↑ | C ↑ | Clear | Clear, better than before | Good | C↑ improved both distance and near |
| L5 | 44.00 | −5.25 | 10.6 | +2.5D | 10.0 ↑↑ | C ↑↑ | Worse | No improvement | — | C too high; dist. degrades, near no lift |

OS Summary

- Power path: −4.75 → −6.00 (over-corrected, back off) → −5.25 (balance point).

- C factor path: 4.0 → 6.0 (improved) → 10.0 (degraded) → high myopia is C-factor-sensitive with an apparent optimum near 6.0.

6. Cross-Comparison

| Dimension | OD — Low Myopia | OS — High Myopia |
| --- | --- | --- |
| C factor sensitivity | Low — 4.0 ↔ 6.0 indistinguishable; only breaks at 1.0 | High — 4.0 → 6.0 improves, 6.0 → 10.0 reverses |
| C factor optimum | ≥ 4.0 sufficient; no need to push higher | ~ 6.0 |
| ADD effect | Core variable; +2.5 D improves near | Not isolated (held at +2.5 D throughout) |
| Power tuning | Not adjusted (−1.25 stable) | Fine-tuning required; narrow window |

7. Overall Findings

(AI-generated synthesis — to be reviewed)

- ADD is the primary driver of near vision. Effect was clearest on the low-myopia eye (R2 → R3).

- C factor sensitivity is power-dependent. Low myopia tolerates a wide C-factor range (4.0–6.0 ≈ equivalent); high myopia shows a narrow optimum near 6.0, with 10.0 actively degrading distance.

- C factor also affects lens cosmesis — the higher the C, the more visible the central ring (barely visible at 1.0 → prominent at 10.0).

- High myopia has a narrower tuning window on both axes — both power over-correction (−6.00) and excessive C factor (10.0) required back-off.

8. Open Questions & Critical Notes

(For review before acting on the findings)

- “Near did not reach expectation” (OD&OS) is a subjective endpoint. Before the next cycle, consider defining a quantitative near-vision threshold (e.g., reading acuity at 40 cm, or a specific J-value) so “meets expectation” is falsifiable.

- ADD was never isolated on the high-myopia track. Whether high myopia responds to ADD the way low myopia does remains untested.

The human is the driver, and AI is the co-driver. @Clark